{"text_format":"1","workbook":"xlsx","n_sheets":2,"sheets":["Berechnung","Datenbank"],"not_rendered":{"charts":2,"images":1,"macros":1,"external_links":0,"defined_names":1,"comments":1,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Berechnung","index":0,"n_rows":34,"n_cols":15,"rows_kept":34,"cut":false,"hidden":false,"merged":[[6,6,6,7]]}
[null,"Bundesamt für Gesundheit, Abt. Strahlenschutz"]
[null,null,"Diagnostische Referenzwerte (DRW)"]
[null,null,"in der Projektionsradiologie",null,null,null,null,null,"BAG-Wegleitung  R-06-04"]
[]
[null,null,"Berechnung der Oberflächendosis am Patienteneintritt"]
[null,null,null,null,null,null,null,null,null,null,1,"Multipulse/HF",null,null,"HF"]
[null,null,"BAG-Bewilligung:",null,null,"A -",null,null,null,null,2,"2-Pulse",null,null,{"$f":"IF(O6=\"2-P\",0.05,IF(O6=\"Multi\",0.1,H11))","v":0.1}]
[null,null,null,null,null,null,null,null,null,null,3,"K-Wert bekannt"]
[null,null,"Röntgenaufnahme:",null,null,1]
[]
[null,null,"Generatortyp:",null,null,null,null,{"$f":"0.1","v":0.1},null,null,1,"Thorax pa",null,0.15]
[null,null,null,null,null,null,null,null,null,null,2,"Thorax lat",null,0.75]
[null,null,"Röhrenspannung: [kV]",null,null,null,null,125,"kV",null,3,"LWS ap",null,7]
[null,null,null,null,null,null,null,null,null,null,4,"LWS lat",null,10]
[null,null,"Strom-Zeitprodukt: [mAs]",null,null,null,null,2,"mAs",null,5,"Becken ap",null,3.5]
[null,null,null,null,null,null,null,null,null,null,6,"Schädel ap / pa",null,2.5]
[null,null,"Fokus-Oberflächendistanz: [m]",null,null,null,null,1.75,"m",null,7,"Schädel lat",null,1.5]
[]
[]
[null,null,"Oberflächendosis am Patienteneintritt:",null,null,null,null,{"$f":"H11*($H$13/100)^2*$H$15*1/($H$17)^2*1.35","v":0.13775510204081634},"mGy"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"H11*($H$13/100)^2*$H$15*1/($H$17)^2*1.35","v":0.13775510204081634}]
[null,null,"Vergleich zum Referenzwert: [100%]",null,null,{"$f":"IF($F$9=1,$N$11,IF($F$9=2,$N$12,IF($F$9=3,$N$13,IF($F$9=4,$N$14,IF($F$9=5,$N$15,IF($F$9=6,$N$16,$N$17))))))&\" \"&\"mGy\"","v":"0.15 mGy"},null,{"$f":"IF(F9=1,(K21/N11)*100,IF(F9=2,(K21/N12)*100,IF(F9=3,(K21/N13)*100,IF(F9=4,(K21/N14)*100,IF(F9=5,(K21/N15)*100,IF(F9=6,(K21/N16)*100,(K21/N17)*100))))))","v":91.83673469387756},"%",null,{"$f":"IF($F$9=1,$N$11,IF($F$9=2,$N$12,IF($F$9=3,$N$13,IF($F$9=4,$N$14,IF($F$9=5,N15,IF($F$9=6,N16,$N$17))))))","v":0.15}]
[null,null,null,null,null,null,null,null,null,null,{"$f":"\"A-\"&G7","v":"A-"}]
[null,null,null,null,null,null,null,null,null,null,{"$f":"IF(E8=1,R12,IF(E8=2,R13,IF(E8=3,R14,IF(E8=4,R15,IF(E8=5,R16,IF(E8=6,R17,IF(E8=7,R18,IF(E8=8,R19,))))))))&IF(E8=9,R20,IF(E8=10,R21,IF(E8=11,R22,IF(E8=12,R23,IF(E8=13,R24,IF(E8=14,R25,IF(E8=15,R26,IF(E8=16,R27,))))))))&IF(E8=17,R28,IF(E8=18,R29,IF(E8=19,R30,IF(E8=20,R31,IF(E8=21,R32,IF(E8=22,R33,IF(E8=23,R34,IF(E8=24,R35,))))))))&IF(E8=25,#REF!,IF(E8=26,R36,))","v":""}]
[]
[]
[null,"V5.1 © 07.2022 BAG"]
[null,null,null,null,null,null,null,null,100]
[null,"Zweckbestimmung:"]
[null,"Die DRWs sind dazu gedacht, die angewandte Praxis bei einem Patientenkollektiv mit dem "]
[null,"korrespondierenden nationalen DRW zu vergleichen. Die DRW sind keine Dosisvorschläge für die"]
[null,"Untersuchung individueller Patienten/-innen und sollten keinesfalls prospektiv verwendet werden. "]
[null,"Das Tool DRWCalc soll zudem die Bewilligungsinhaber/-innen bei der Dokumentation der angewandten "]
[null,"Praxis unterstützen."]

{"sheet":"Datenbank","index":1,"n_rows":2,"n_cols":10,"rows_kept":2,"cut":false,"hidden":false,"merged":[[0,2,0,3]]}
["Datum","Region","Generator\n(K-Konstante)",null,"kV","mAs","FOD\n[m]","Ref. Dosis\n[mGy]","Dosis\n[mGy]","Vergleich\n[%]"]
["a"]